APPENDIX E:

MATLAB CODE TO DETERMINE THE RLC VALUES FOR THE CIRCUIT MODEL

In this appendix, the MATLAB code that was used to calculate the RLC values from the S11 values measured by the network analyzer is provided.  The calculations used by the code are described in Chapter 5.

# Main Program

%This is a script that will determine the proper frequency

%to use as the normalization value in the linear

%extrapolation.  The frequency selected is the frequency

%corresponding to the maximum value of Real(1/Zout) of the

%transducer.  The S11 parameter for each transducer were

%measured using a Network Analyzer.  The program also

%determines the RLC equivalent circuit for each transducer.

clear all;

close all;

j=sqrt(-1);

%Enter the data file to be read in

dat='/dunn/bigelow/Transducer_impedance/DATA06num.txt'

Zo=50; %Impedance of line used by NA to find S11.

N=7; %Set number of points on either side of peak to find

%fit.

%Read in the data

[S11_re,S11_im,freq]=read_S11(dat);

S11=S11_re + j*S11_im;

freq=freq*1e-8; %Adjust to improve fit calculation

%Determine the Input Impedance at each frequency

Zin=Zo*(1+S11)./(1-S11);

Yin=1./Zin;

%Find the maximum in real(Yin)

[m,n]=max(real(Yin));

%Determine polynomial coefficients (degree = 8)

P=polyfit(freq(n-N:n+N),(Yin(n-N:n+N)),8);

%Extract region over which extrapolation will be done.

f=freq(n-N):((freq(n+N)-freq(n-N))/200):freq(n+N);

y=polyval(P,f);

ye=100*abs(polyval(P,freq(n-N:n+N)) - …

Yin(n-N:n+N))./abs(Yin(n-N:n+N));

max(ye)

%Plot the Real(Yin)

figure(1)

clf

plot(freq*100,real(Yin))

hold

plot(f*100,real(y),'r-.')

grid

xlabel('Frequency (MHz)')

ylabel('Real(Yin) (Mhos)')

%Determine Resonance Frequency

[m2,n2]=max(real(y));

freq_res=f(n2)*100

%Determine R1

R1=1/m2

%Determine Co

Co=imag(y(n2))/(2*pi*f(n2)*1e8)

%Find X1

w=1e8*f*2*pi;

wo=2*pi*1e8*f(n2);

X1=imag(1./(y-j*w*Co));

%Determine C1

C1=0;

cnt=0;

for wi=1:length(X1)

if abs(X1(wi))>1e-9

val=((w(wi)/wo)^2 - 1)/(w(wi)*X1(wi));

C1=C1+val;

cnt=cnt+1;

end

end

C1=C1/cnt

%Determine L1

L1=1/(C1*wo^2)

file	= input('Enter a filename to save the transducer information: ','s');

if ~isempty(file)

save(file,'freq_res','R1','Co','L1','C1');

'Saved Resonance Frequency, R1, Co, L1, and C1 !!!!!'

end

Read in S11 Values

function [S11_re,S11_im,f]=read_S11(dat);

%This is a MATLAB function that reads in the S11 values from

%the Network Analyzer.

%Inputs

%    dat = data file

%Output

%    S11_re = real part of S11

%    S11_im = imag part of S11

%    f = frequency of measurement point

%***********************************************************

fid	= fopen(dat, 'r');

fstart=fscanf(fid,'%f',[1,1]);

fend=fscanf(fid,'%f',[1,1]);

fnum=fscanf(fid,'%i',[1,1]);

for si=1:fnum

S11_re(si)=fscanf(fid,'%f',[1,1]);

hld=fscanf(fid,'%c',[1,1]);

S11_im(si)=fscanf(fid,'%f',[1,1]);

f(si)=fstart+(si-1)*(fend-fstart)/(fnum-1);

end

fclose(fid);